Flow Cytometry Core Facility – Canto II V4.2 Email completed form to fccf@salk.edu for scheduling access

Conditions and Acknowledgement form for use: Read also Usage Guide on Salkland  salk.edu/fccfaccess > Help Files

- FCCF has analyzers in 2 places: “Parker Room” (A3/LSRII-2/MACSQuant VYB/Canto) or “STEM” (LSRII-1) that can be booked concurrently (are more than 6 feet apart).

- Buffer time between users: 30min between users to allow for clean up and unexpected issues

- Instrument operation is subject to new usage instructions available on Salkland: salk.edu/fccfaccess > Help Files

- Your first autonomous session: Plan to come M-Th during the day, and ask staff when you should book your session (we will find a good time for you to book, in case you need help). Book extra time to learn how run the QC.

- Parker Room occupancy: Maintain 6ft distance while in room (priority given to instrument operators).

- Only one person at an instrument unless pre-approved by core. Any time you bring a co-worker, please indicate their attendance/name on the Log sheet (for contact tracing).

- Always check the calendar before going to the lab: be 100% aware of who is working especially if you are jumping in on short notice ahead of a user – obtain their permission first if there is any chance you are running late and could be present when they come to use the instrument.

- Changing to another instrument: Make a new calendar booking on the instrument you are using for contact tracing (inform the core if you need cancellation help on the original booking).

- Analyzers are running autoclaved MQ water (instead of PBS) as sheath at this time.

- You may be asked to execute an extended cleaning protocol if the post use QC shows it is needed.

- No drop ins, bookings/appointments only: You must reserve a session on Infinity BEFORE using the machine. Your first autonomous session: Plan to come M-Th during the day, and book extra time to learn how run the QC; ask staff when you should book your session in case you need help.

- COMMUNICATE with each other by email, especially if you are delayed on an instrument.

- Enable wifi calling to use cell phone in the basement areas (FCCF or STEM).

- Access requires protocol review by FCCF (for inclusion of DNAse I: which steps/how much): Submit your protocol for review along with this form.

- Unclogging the instrument: You will need to demonstrate to staff that you know how to insert the cleaning stylus wire for both types of clog scenarios. Full obstruction: no events displayed, nothing taken up from tube and Prime button does not generate bubbles; partial obstruction: weird looking scatter pattern with sample and with beads, event rate fluctuates (clumpy sample).

| Name: | Lab: | Cell phone: |  |
| --- | --- | --- | --- |
| My protocol has been approved by FCCF (steps indicating use of DNASe I: e.g./how much) |  |  |  |
| I understand that the instrument is running MQ water as sheath fluid instead of PBS |  |  |  |
| I have already located all the instrument Help Files in Salkland or Infinity and read them |  |  |  |
| I know where to locate more: Bleach (by door in FCCF), CS&T beads in STEM, Parker Room (or in FCCF), and MQ water sheath carboys (if not nearby, autoclaved carboys are delivered by Salk Glass-wash staff to the corridor area immediately outside FCCF’s main door) |  |  |  |
| I will put rinsed empty bleach bottles into recycling, and change the red bio trash bag if full |  |  |  |
| I will log into Diva under Calibration (pw = orangecst) to look at CST bead profiles and make PDFs |  |  |  |
| I will check instrument cleanliness first, before running beads, looking for low event rate (fluctuating as low as 0-20evt/s when very clean, less than 50-100 evts/s) with freshly filtered water. |  |  |  |
| I will change the syringe filter if I see high event rate for water, and will re-check before using to make up beads |  |  |  |
| If I am the first user of the day, I will execute the automated calibration and export a PDF to send to fccf@salk.edu |  |  |  |
| I will produce a “before” bead profile PDF if (1) I’m the first user of the day or (2) a user immediately running before me on the same day has not left an acceptable “after” file in the Calibration login (that will suffice as “before” data for me). |  |  |  |
| I have accessed Help Files in Salkland or Infinity and read the Usage Guide; I will comply with providing: |  |  |  |
| (i)“5min cleaning” as described every 30min use (if samples prove difficult to clean up after) |  |  |  |
| (ii) Recording cleaning files from the 5min cleaning for (a) bleach, (b) rinse and (c) water |  |  |  |
| (iii) Checking after use for cleanliness according to instructions, or staff directions |  |  |  |
| (iv) Submitting screengrab proof (eg Diva Browser with time stamps) |  |  |  |
| I will read and will perform every step and sign off the clipboard after use: |  |  |  |
| (i) Tank (Sheath & Waste) maintenance |  |  |  |
| (ii) Data backup I will delete my previous session’s data |  |  |  |
| (iv) Instrument OFF I will take a photo showing the GREEN POWER BUTTON (OFF) |  |  |  |
| (v) I understand that ONLY A PHOTO OF THE POWER BUTTON IS ACCEPTABLE for (iv) |  |  |  |
| (vI) I will take a photo of the completed/signed clipboard |  |  |  |
| I will restart the PC before I leave (restart, NOT shut down) |  |  |  |
| I will submit all files and photos to fccf@salk.edu |  |  |  |
| (i) Beads before/after, CST report in PDF format as appropriate |  |  |  |
| (ii) Photo/screengrab of Diva browser showing cleaning files with time stamps |  |  |  |
| (iii) Photo of signed clipboard with instrument power button to show how I left it |  |  |  |
| I will perform corrective maintenance (long cleaning) under staff directions if needed |  |  |  |
| I have looked at the examples on the pages following this form, and understand exactly which files/photos to send immediately after my session to fccf@salk.edu |  |  |  |
| I understand that “Duplicate without Data” is NOT TO BE USED as it can break the DIVA database; I will ask FCCF staff how to correctly re-use instrument settings |  |  |  |
| Troubleshooting instrument issues: I understand how to use the Cytometer Status (Tera Term) desktop icon to open a window that will allow me to copy computer code if needed (such as error codes) |  |  |  |
| I understand what a full vs a partial clog looks like and have demonstrated to staff I can insert a cleaning wire |  |  |  |
| New instructions may require another form to acknowledge my reading and understanding |  |  |  |

Good Examples of What You Are REQUIRED to Send to FCCF@salk.edu DIRECTLY AFTER Using an Instrument.

Emailing items 3, 4, and 5 (unless specifically directed otherwise) is a condition of your access to the core’s autonomous use instrumentation. Item 1 is required if the user before you does not have an acceptable “before” beads in the Calibration account. Item 2 is required if you are the first user of the day.

- CST Beads Profile PDF (Pre-running samples): required if the user before you does not have an acceptable “before” beads in the Calibration account

- CS&T Report PDF (2 pages): required if you are the first user of the day

- Cleaning Files: Screen grab with Time Stamps

- CST Beads PDF (post-running samples and post final round of cleaning):

- Clipboard checklist items, signed (legible), with power button OFF in the background: