Flow Cytometry Core Facility – A3 V2 Email completed form to fccf@salk.edu for scheduling access

Conditions and Acknowledgement form for use. Read also Usage Guide on Salkland salk.edu/fccfaccess > Help Files

- FCCF has analyzers in 2 places: “Parker Room” (A3/LSRII-2/MACSQuant/Canto) or “STEM” (LSRII-1) that can be booked concurrently (are more than 6 feet apart).

- Buffer time between users: 30min between users to allow for clean up and unexpected issues

- Instrument operation is subject to new usage instructions available on Salkland: salk.edu/fccfaccess > Help Files

- Your first autonomous session: Plan to come M-Th during the day, and ask staff when you should book your session (we will find a good time for you to book, in case you need help). Book extra time to learn how run the QC.

- Parker Room occupancy: Maintain 6ft distance while in room (priority given to instrument operators).

- Only one person at an instrument unless pre-approved by core. Any time you bring a co-worker, please indicate their attendance/name on the Logsheet (for contact tracing).

- Always check the calendar before going to the lab: be 100% aware of who is working especially if you are jumping in on short notice ahead of a user – obtain their permission first if there is any chance you are running late and could be present when they come to use the instrument.

- Changing to another instrument: Make a new calendar booking on the instrument you are using for contact tracing (inform the core if you need cancellation help on the original booking).

- Analyzers are running autoclaved MQ water (instead of PBS) as sheath at this time

- You may be asked to execute an extended cleaning protocol if the post use QC shows it is needed

- No drop ins, bookings/appointments only: You must reserve a session on Infinity BEFORE using the machine. Your first autonomous session: Plan to come M-Th during the day, and book extra time to learn how run the QC; ask staff when you should book your session in case you need help.

- COMMUNICATE with each other by email, especially if you are delayed on an instrument

- Enable wifi calling to use cell phone in the basement areas (FCCF or STEM)

- Access requires protocol review by FCCF (for inclusion of DNAse I: which steps/how much). Please send over your protocol for review along with this form.

- Unclogging the instrument: You will need to demonstrate to staff that you know how to insert the cleaning stylus wire for both types of clog scenarios. Full obstruction: no events displayed, nothing taken up from tube and Prime button does not generate bubbles; partial obstruction: weird looking scatter pattern with sample and with beads, event rate fluctuates (clumpy sample).

- Sign up for A3-Users mailing list for updates on instrument protocols, repairs, and any other relevant information regarding the usage and health of the A3. You can do this by visiting http://lists.salk.edu/mailman/listinfo/a3-users , filling out the information in the Subscribing to A3-Users section, then hitting the “subscribe” button at the bottom of the section.

| Name: | Lab: | Cell phone: |  |
| --- | --- | --- | --- |
| My protocol has been approved by FCCF (steps indicating use of DNASe I: e.g. which step, and how much) |  |  |  |
| I understand the A3 runs on buffer with surfactant and will fill tank using the labeled carboys located by the A3 |  |  |  |
| I understand the A3 runs on buffer with surfactant and will check with FCCF staff regarding sensitive live cell assays. |  |  |  |
| I will perform tank maintenance if there is 3L or less sheath / 7L or more waste. |  |  |  |
| I am aware of the two audible alarms (sheath and waste) and troubleshooting steps if the waste alarms erroneously |  |  |  |
| I know how to locate instrument related documents (eg extra troubleshooting SOPS ) on Salkland at salk.edu/fccfaccess > Help Files (a desktop link is provided on the A3 workstation). |  |  |  |
| I know where to locate more: Bleach (by door in FCCF), CS&T beads in STEM, Parker Room (or in FCCF), and MQ water sheath carboys (if not nearby, autoclaved carboys are delivered by Salk Glass-wash staff to the corridor area immediately outside FCCF’s main door) |  |  |  |
| I will put rinsed empty bleach bottles into recycling, and change the red bio trash bag if full |  |  |  |
| I understand how, and will purge the sheath filter of air before I commence usage |  |  |  |
| I understand how to trouble-shoot for missing lasers according to instructions (e.g. in Help Files): |  |  |  |
| (i) FCCF staff have already set up Coherent Connect (CC) for me during training to show only the UV laser |  |  |  |
| (ii) I am aware that a troubleshooting procedure for Coherent Connect will be located in Help Files for the A3 |  |  |  |
| (iii) I understand that I can check that a laser is ON by running beads in the QC template (Calibration login) if there are issues with a laser I do not need, I CAN proceed to use the instrument |  |  |  |
| (v) I will exit Coherant Connect after turning ON/OFF to avoid lasers turned off randomly |  |  |  |
| (vi) I am able to turn on the UV laser (it requires both the power slider as well as the START button in CC!!) |  |  |  |
| I will log into Diva under Calibration (pw =bluecst) to look at CST bead profiles and make PDFs |  |  |  |
| I will check instrument cleanliness first, before running beads, looking for low event rate (fluctuating as low as 0-20evt/s when very clean, less than 50-100 evts/s) with freshly filtered water. |  |  |  |
| I will change the syringe filter if I see high event rate for water, and will re-check before using to make up beads |  |  |  |
| If I am the first user of the day, I will execute the automated calibration and export a PDF to send to fccf@salk.edu |  |  |  |
| I will produce a “before” bead profile PDF if (1) I’m the first user of the day or (2) a user immediately running before me on the same day has not left an acceptable “after” file in the Calibration log in (that will suffice as “before” bead data for me). I will always produce an “after” bead profile following cleaning up, to demonstrate the LSRII is clean. |  |  |  |
| I will run the automated CS&T program if I am the first user (or if I am troubleshooting a problem), and: |  |  |  |
| (i) I will NOT run the program UNLESS ALL 5 LASERS ARE ON (detectors have peaks in all histograms) |  |  |  |
| (ii) I am aware that the program will time out (and give error messages) if the beads are too dilute |  |  |  |
| (iii) Beads are $$, I will put the diluted/labelled bead tube back in the fridge so they can possibly be used again |  |  |  |
| I understand that unstable/shifting signals (eg green laser on LSRII-1) is something I can probably fix by: |  |  |  |
| (i) First purging any air from the sheath filter |  |  |  |
| (ii) Second, removing at the flow cell: run an empty tube as soon as bubbles appear in the waste line (tubing that runs to the top of the waste tank) run water 5min |  |  |  |
| (iii) Executing the automated CS&T to calibrate the laser delay values now that air has been purged |  |  |  |
| I have accessed Help Files in Infinity and read the Usage Guide; I will comply with: |  |  |  |
| (i)“5min cleaning” as described every 30min use (or less) if my samples prove difficult to clean up after |  |  |  |
| (ii) Recording cleaning files from the 5min cleaning for (a) bleach, (b) rinse and (c) water. |  |  |  |
| - Checking after use for cleanliness according to instructions, or staff directions |  |  |  |
| (iii) Submitting screengrab proof (eg Diva Browser with time stamps) |  |  |  |
| I will read and will perform every step and sign off the clipboard after use: |  |  |  |
| (i) Tank (Sheath & Waste) maintenance |  |  |  |
| (ii) Data backup I will delete my previous session’s data to help maintain Diva database at less than 4gb |  |  |  |
| (iv) Instrument OFF I will take photo of fluidics panel, together with clipboard |  |  |  |
| (v) I understand that ONLY A PHOTO OF THE FLUIDICS PANEL IS ACCEPTABLE for (iv) |  |  |  |
| I will restart the PC before I leave (restart, NOT shut down) |  |  |  |
| MANDATORY: Before leaving, I will send a Photo of signed clipboard with instrument fluidics panel showing how I left it (eg OFF) to the next user on the schedule (if there is one) and fccf@salk.edu. |  |  |  |
| I will perform corrective maintenance (long cleaning) under staff directions if needed |  |  |  |
| I understand that “Duplicate without Data” is NOT TO BE USED as it can break DIVA database; I will ask FCCF staff how to correctly re-use instrument settings (i.e. I will load templates correctly, from the Experiment Menu) |  |  |  |
| I understand that Diva is periodically reinstalled (due to database issues) and I may need to make myself a new account using the provided instructions by the LSRII at my next returning session if I cannot log in to Diva |  |  |  |
| I understand what a full vs a partial clog looks like and have demonstrated to staff I can insert a cleaning wire |  |  |  |
| I shall ensure a tube with NO MORE than 1mL of H2O is installed on the instrument before putting in STANDBY or OFF |  |  |  |
| New instructions may require another form to acknowledge my reading and understanding |  |  |  |

Good Examples of What You Are REQUIRED to Send to FCCF@salk.edu DIRECTLY AFTER Using an Instrument.

Emailing items 3, 4, and 5 (unless specifically directed otherwise) is a condition of your access to the core’s autonomous use instrumentation. Item 1 is required if the user before you does not have an acceptable “before” beads in the Calibration account. Item 2 is required if you are the first user of the day.

- CST Beads Profile PDF (Pre-running samples): required if the user before you does not have an acceptable “before” beads in the Calibration account.

- CS&T Report PDF (2 pages): required if you are the first user of the day.

- Cleaning Files: Screen grab with Time Stamps

- CST Beads PDF (post-running samples, post final round of cleaning):

- Clipboard checklist items, signed (legible), with fluidics panel OFF in the background: